Reference, Maintenance 

& User's Guide for 

the Knee Ligament Arthrometer®
MEDmetric®
Knee Ligament ARTHROMETER®
Models KT1000™ and KT2000™

Dale M. Daniel, M.D.

Associate Clinical Professor of Orthopedic Surgery

University of California, San Diego, School of Medicine

Staff Orthopedic Surgeon

Kaiser Permanente, San Diego
First Edition
May 1993.
Revised December 1993; May 1994; July 1994; October 1996, December 1998, March 2005

MEDmetric® Corporation, 7542 Trade Street, San Diego, California 92121-2412
Copyright

© 1993 by MEDmetric®  Corporation. All rights reserved. This book is protected by copyright. No part of it may be
reproduced, stored in a retrieval system, or transmitted in any form or by any means, electronic, mechanical, photocopying or recording, or otherwise,  for any purpose other than the purchaser’s personal use, without prior written permission of MEDmetric® Corporation.

Intellectual Property
Arthrometer, KT1000™, KT1000/S,™ ,KT2000™ and MEDmetric are proprietary marks owned by MEDmetric® Corporation. The apparatus and methods of the KT1000™ and KT2000™ are protected by U.S. Patent 4,583,555.


Made in the United States of America

MEDmetric®   Corporation

 Reference, Maintenance & User’s Guide For The Knee Ligament Arthrometer.



MEDmetric® Corporation has designated the following representative within the European community as its Authorized European Representative pursuant  to Article XIV of the Medical Device Directive 93/42/EEC 

Robin Humble

68L Chapel Mews; Crewe Road

Alsager; Staffordshire, ST72HA

United Kingdom

Tel: (44) 1270883342

Fax: (44) 1270875622

BLANK PAGE

Contents
Thank You
1

Introduction
2

Satisfaction Assurance
4

The ARTHROMETER® System
5

Glossary of Terminology
7

Part 1: User’s Guide
Overview
9
The Testing Sequence
9
Developing a Consistent Technique
10
Biomechanics of the Knee
11
Coupled Motion and Constraint
12
Flexion and the Ligaments
13
Getting Started
15
ARTHROMETER® Accessories and Their Use
15
Setting Up for PCL Tests
16
Setting Up for ACL Tests
17
The Testing Reference Position
19
Patient Relaxation Checklist and Other Considerations
20
Setting Up the X-Y Plotter (Model  KT2000™)
21
Setting Up the X-Y Plotter (Recent Model KT2000™)
22
Testing Technique
23
Pull-Push Testing Cycle (KT1000™)
23
Pull-Push Testing Cycle (KT2000™)
24
Push-Pull Testing Cycle (KT2000™)
25
Testing Cruciate Ligaments
26
90° Quadriceps Active Test (Visual)
26
PCL Assessment: Tests at the Quadriceps Neutral Angle 
27
ACL Assessment: Passive Displacement Test
29
ACL Assessment: The Compliance Index
30


ACL Assessment: Quadriceps Active Test at 30°
31
ACL Assessment: Manual Maximum Displacement Test
32
Data Interpretation
33
Sources of Measurement Error
33
References
34

Documenting Knee Stability in the O.R...................................................................................................35

Part 2: Maintenance
Maintenance and Troubleshooting
36


Force Tone Checks
37
Posterior Force Tones
37
Anterior Force Tones
37
Checking Accuracy
38
Checking Linearity
39
Procedures for the X-Y Plotter
40

Limited Warranty
41

Terms and Conditions of Sale
42

Appendices
Appendix A
43


Side-to-Side Differences: ACL Deficit Patients Compared to Normal Subjects
44
Test-Retest Variations
44
Measurement Accuracy (Cadaver Studies)
45
Appendix B
46
Electronic Specifications for the KT2000™
46
Interface Cord Color Code
46
KT1000™/KT2000™ Patient Evaluation Form Sample
47

KT2000™ X-Y Plotter Patient Evaluation Form Sample
48

Thank You
Thank you for acquiring the MEDmetric® Knee Ligament ARTHROMETER®. You have selected the first available, most used, and most accurate knee ligament testing system manufactured, a claim supported by the many published scientific papers and articles produced using data collected with the  KT1000™ and KT2000™.

The KT1000™, first introduced in 1982, is the product of research and development beginning prior to 1979. Orthopedic surgeon Dale M. Daniel, M.D., devised schemes for instrumenting  knee articulation and, with the collaboration of MEDmetric® founder Lawrence L. Malcom, Ph.D., developed three distinct prototypical models which evolved finally to the portable form of the instrument as it exists today.

Design improvements are continually incorporated which enhance the accuracy, reliability and durability of the KT devices. As significant improvements or new features are developed, you will be notified and given the opportunity to upgrade your instrument. Your KT device need never be obsolete.

Although the ARTHROMETER® is easy to use and examinations may be performed quickly, a reliable and reproducible technique, requiring study and practice, must be developed. A modest commitment of  time is needed to acquire a further understanding of knee mechanics and to develop a good testing procedure. The support materials supplied are intended to assist  examiners in their efforts. Please review these carefully. Regional Knee Ligament ARTHROMETER® workshops are conducted from time to time; we encourage you to attend if possible. Alternatively, we suggest that you network with KT device users in your area and spend time with examiners who have a developed technique. MEDmetric® Corporation phone support is also available for your convenience.

We are very interested in your comments regarding the Knee Ligament ARTHROMETER® and its use. Your observations may guide us in producing a design enhancement which will improve performance of the instrument. Please address all correspondence to:

MEDmetric® Corporation

7542 Trade Street

San Diego, California 92121-2412

Phone: (619) 536-9122 or (800) 995-6066;

Fax: (619) 536-9303; E-Mail: info@medmetric.com

[image: image1.png]V2%




Dick Watkins

President

Introduction


The KT1000™ was developed to provide objective measurement of the sagittal plane motions of the tibia relative to the femur. This motion, sometimes referred to as drawer motion, occurs when an examiner applies force to the lower limb or when the muscles of the quadriceps are contracted. Both the KT1000™ and KT2000™ models of the Knee Ligament ARTHROMETER® accurate, easy-to-use instruments for the clinical assessment of ACL and PCL integrity.

· There is no anesthesia risk.

· The patient undergoes no x-ray exposure.

· Due to patient comfort, there is minimal muscle guarding.

· It is inexpensive when compared to x-ray, MRI or CT scan testing.

Since the Knee Ligament ARTHROMETER® was designed for the orthopedic practitioner, the 

instrument allows the examiner easily to screen acute knee injuries and quantitatively document 

anterior-posterior knee measurability in the office or clinical setting.

Ease of Testing Large Patients
The Knee Ligament ARTHROMETER® was designed to overcome problems commonly encountered when attempting to perform a 25° displacement measurement on a large, well muscled, or obese patient. This was accomplished by supporting both things with a firm, comfortable platform placed proximal to the popliteal space. This helps keep the patient's knee flexion angle constant throughout the test. A foot support accessory, supplied with all ARTHROMETERs®, positions the feet symmetrically, allowing optimal leg positioning for the test while reducing external rotation of the tibia.

Contributes to Patient Relaxation
The patient seems to relax and in general react favorably to an ACL stability examination done with an ARTHROMETER®. The flexion angle of the test, which is between 20° and 35°, is usually the most comfortable position for an acute knee injury. A thigh strap (provided) controls external hip rotation and offers support that encourages patient relaxation. Since both legs are lying on a stable thigh support platform, the patient can relax and will exhibit less apprehension during testing. 

Improvements, Upgrades and Retrofits
Over the years, MEDmetric® has modified the Knee Ligament ARTHROMETER® and its 

accessories to improve performance. The information in this guide reflects the KT1000™ and 

KT2000™ as of October 1996. Among the differences found in earlier model KTs are:

· Hall Effect Retrofit (ARTHROMETERs® with serial numbers below 186). The earliest KT models required periodic adjustment or recalibration of force tones. The Hall Effect Retrofit eliminated the need for scheduled service.

· Thigh Strap (KT serial number 200). The strap was introduced to resist external limb rotation during  ACL testing at 20° to 35° of flexion.

· 30 lb. Force Tone. The original KT1000™ design incorporated two anterior force levels: 15 lb. and 20 lb. A 30 lb. force tone was later incorporated to add test sensitivity for well muscled or large patients.

· Adjustable Thigh Support. The thigh support for current production ARTHROMETERs® is adjustable in three steps for short, medium and tall subjects. The new thigh support is MRI safe since it has no metal components.

· Posterior Force Tones. KT2000™ (serial number 455 and higher) and the newest KT1000™  units emit force tones when 15, 20, and 30 pounds of force is applied posteriorly. Earlier KT1000™ devices featured posterior force tones at 15 and 20 pounds; previous KT2000™ units enunciated a tone at 20 pounds of posterior force only.

· Instructional Video. The video accompanying the ARTHROMETER® has been updated (© 1991).

· X-Y Plotter. The KT2000™ X-Y plotter has been upgraded.

· Force Tone Shutdown. When batteries are weak, new KT2000™ units (Serial Number 327 and higher) cease enunciating force tones. This is an indication that batteries must be changed in order to maintain plotter/computer output accuracy.

Factory installed retrofits and upgrades are available for all Knee Ligament  ARTHROMETERs®. To find the serial number of your KT, examine the metal plate affixed to the bottom of your Knee Ligament ARTHROMETER®.

NOTE: ARTHROMETER®, KT1000™, KT2000™ and MEDmetric® are trademarks of MEDmetric®  Corporation.  Utility and methods employed in the design and use of the MEDmetric® Knee Ligament ARTHROMETER® (Models KT1000™ and KT2000™) are protected by U.S. Patent 4,583,555. All rights strictly reserved.

Satisfaction Assurance


MEDmetric® Corporation assures customer satisfaction. If for any reason during the first thirty (30)  days following delivery of your ARTHROMETER® you are dissatisfied with it, you may return your instrument in "as new" condition for full credit of the price paid or pledged.

We ask only that you provide us with the cause of your dissatisfaction in writing and that you inform us of your intention prior to return of the instrument.

Shipping charges are the responsibility of the purchaser. In the event of returned equipment, your 

account will be credited with the purchase price and applicable sales taxes only. You will be billed for the shipping charges.

We feel that thirty (30) days of use is adequate to determine your acceptance of the instrument. 

Therefore, no extensions of the thirty (30) day satisfaction guarantee will be made, except by express written authorization.

[image: image2.png]V2%




Dick Watkins

President

The ARTHROMETER® System
[image: image3.png]


    

	A - Battery Compartment Door
	E - Displacement Dial

	B - Distal Velcro Strap
	F - Patella Sensor Adjustment Dial

	C - Proximal Velcro Strap
	G - Patella Sensor Pad

	D - Force Handle
	H - Joint Line Arrow

	
	I  -  Tibia Sensor Pad


· Knee Ligament ARTHROMETER® anterior/posterior tibial displacement measuring instrument. 

· Adjustable Thigh Support Platform. Raises thighs 11 cm - 17.5 cm from table surface.

· Thigh Strap. To stabilize limbs.

· Foot Support
· Goniometer
· Carrying Case

· User's Guide

· Instructional Video

· Scripted 35 mm Slide Presentation

 Research Results Reprints: 

 The Accuracy and Reproducibility of the KT1000™

 The Active Drawer Test Instrumented Measurement of Acute ACL Disruption

 Knee Stability Measurement Using the KT1000™ Knee Ligament ARTHROMETER®
 The Quadriceps Active Test
· Patient Evaluation Forms
KT2000™ X-Y plotters: (Fig. 1) The model used on KT2000 ™ devices made between 1992 and 1996. (Fig. 2) Current version.



(Figure 1)

A - Power

     
B - Off-Chart-Pen Switch

     
C - "X" Adjustment Knob

     
D - "Y" Adjustment Knob

     
E - Pen


(Figure 2)

A - Power Pen Up/Down Lever

B - Paper Clamp

C - Pen

D - "X" Adjustment Knob

E - "Y" Adjustment Knob
Supplemental KT2000™ Materials:

· X-Y Plotter

· Coiled ARTHROMETER® Plotter Cord

· Graph Evaluation Template

· X-Y Plotter Carrying Case (Optional)

· Plotter Pens

· Patient Evaluation Forms & Graph Paper
Glossary of Terminology


Active displacement. Tibial displacement when force is provided by the patient when contracting the quadriceps muscles. Referred to as the "active drawer" in some literature.

ARTHROMETER®. MEDmetric® tibio-femoral joint displacement measuring instrument in two models, the KT1000™ and KT2000™.

compliance index. Tibial displacement occurring between two displacement forces during a continuous testing cycle.

compliance index difference (CID). The relative side-to-side difference in each patient's compliance index measurement.

corrected anterior displacement for the PCL injured knee. The passive anterior displacement minus the quadriceps active displacement when both are measured relative to the testing reference position at the quadriceps neutral angle.

corrected posterior displacement for the PCL injured knee. The passive posterior displacement plus the quadriceps active displacement measured relative to the testing reference position at the quadriceps neutral angle.

foot support. A support placed under both feet to establish symmetrical axial knee position before each test.

involved knee (I). The knee with the current injury or complaint of instability.

manual maximum displacement. The maximum anterior tibial displacement which can be produced by a high manual force applied to the proximal calf. Referred to as the "manual maximum drawer" in some literature.

non-involved knee (N). The knee of the patient opposite to the involved knee. If the non-involved knee is normal, measurements of the involved knee should be compared to the non-involved knee (I-N).

passive displacement. Tibial displacement when force is provided by the examiner. Referred to as the `` passive drawer" in some literature.

quadriceps active displacement. Tibial displacement caused by active contraction of the patient's quadriceps muscle.

quadriceps neutral angle (QNA). The flexion angle in a normal knee where there is no tibial movement either anterior or posterior during an active quadriceps contraction. The quadriceps neutral angle is approximately 70°, although the angle varies from patient to patient.

side-to-side difference. The relative measurement difference between a patient's two knees.

testing reference position. Tibio-femoral position in a relaxed patient after a 20 lb. posterior push has been applied and released.

thigh support. A platform placed under both thighs to maintain symmetrical knee flexion.

tibial sag screen. A quadriceps active displacement test at 90° of flexion. A normal knee moves slightly posterior .5 to 2 mm.

total anterior-posterior displacement. Total displacement resulting from a 20 lb. posterior push and a 30 lb. anterior pull.

Overview

This User's Guide provides instruction in the operation and maintenance of the  MEDmetric® Knee Ligament ARTHROMETER®. Documentation for the ARTHROMETER® consists of this book, which describes the ARTHROMETER® system and the ways you can work with it; a video tape,  Instrumented Drawer Testing; a scripted slide show; and reprints and summaries of some of the significant studies and papers which have made use of the KT1000™ and KT2000™.

The Knee Ligament ARTHROMETER® measures the relative motion between a patellar sensor pad and a tibial sensor pad. The ARTHROMETER® is secured to a patient's lower leg. When the force handle is pulled, the tibia displaces anteriorly and the KT indicates this motion in millimeters. When the handle is pushed, the tibia displaces posteriorly and the KT indicates posterior displacement. When the handle is released following a pull-push sequence, the tibia will return to the testing reference position.


[image: image4.wmf]

Quad Active







Manual Maximum







20º - 35º ACL

Passive Displacement





Pass







Measure for

Complience

Index Only



Quad Active







Passive Displacement







Quadriceps Neutral

Angle -70º





Fail

Measure for both ACL & PCL





(Quadriceps Active)

Instrumented





(Quadriceps Active)

Visual





Screen for 90º PCL








In a normal knee, anterior force results in anterior displacement and internal rotation of the tibia; posterior force results in posterior displacement and external rotation of the tibia. These A-P displacements and rotations of the tibia are coupled behavior in the knee joint. When cruciate ligaments are torn, both displacement and rotation patterns of the tibia change. Cruciate ligament deficiency can be detected in most knees by comparative measurement of tibial displacement alone.

The Testing Sequence
A Knee Ligament ARTHROMETER® examination sequence consists of tests to determine the integrity of the posterior cruciate ligament followed by tests to determine the integrity of the anterior cruciate ligament. With a relaxed, cooperative patient, an examiner experienced in the operation of the ARTHROMETER® can perform the entire series of exams in less than 10 minutes.

Generally, ARTHROMETER® exams consist of testing the non-involved (or "normal") knee and repeating each test on the involved (or injured) knee. A variation in measurements between the non-involved and involved knees can indicate cruciate ligament disruption.

PCL Screen. The test series begins with a visual quadriceps active test with the patient's knee flexed to 90° followed by an Instrumented exam at the same angle of flexion. When no posterior cruciate ligament disruption is indicated, the examiner proceeds to the ACL test cycle; when tests indicate a PCL disruption, tests at the quadriceps neutral angle are performed.

Testing for ACL and PCL in the PCL Deficient Knee. To quantify PCL disruption, you must first establish your patient's quadriceps neutral angle (QNA). This angle varies from patient to patient, but is approximately 70°. Once you have determined your patient's QNA, active tests (those involving patient muscle contraction) and passive tests (where the examiner applies pressure) will confirm and quantify PCL disruption. Passive displacement tests at the QNA can quantify ACL as well as PCL disruption.

ACL Testing. Tests for anterior cruciate ligament disruption are performed at between 20° and 35° of knee flexion. Three tests are performed at this angle: the passive displacement test (to test for ACL deficiencies); the manual maximum displacement test (endpoint measurement); and the quadriceps active test (to measure tibial displacement with a quadriceps muscle contraction). In addition, a "compliance index" (the difference between readings at different force levels) can quantify the feeling of a "soft endpoint" which is often observed after an ACL disruption. Note: Only the compliance index can be used for ACL testing in the PCL deficient knee.

Developing a Consistent Technique
The Knee Ligament ARTHROMETER® provides an effective means of quantifying knee laxity, and is a valuable means of monitoring recovery in a patient undergoing therapy. The effectiveness of  ARTHROMETER® testing is directly dependent on the development of testing technique. Ensuring that variables such as knee flexion angle and force vector angle are reproduced and controlled will enhance accuracy of measurement.

When retesting patients, examiners should refer to previous Patient Evaluation Forms and re-create the angles of flexion and heel position as recorded. The goniometer, an integral part of the ARTHROMETER® system, is designed for this.  In the field, there ate two methods of using the goniometer. The first (and more precise) is to align the pivot point along the joint line, adjusting the arms of the goniometer with the femur and tibia.  The second method is to place the goniometer on top of the patella, resting one arm on the tibial spine and the other on the anterior aspect of the thigh. Either method will work, but the same procedure must be employed by all examiners within a group to ensure valid test-retest results.

Other variables may affect results, including patient relaxation, ARTHROMETER® placement, consistent patella pad pressure and speed and vector of force application.. The development of a consistent and standardized testing will lead to precise and reproducible anterior-posterior displacement measurements.

Biomechanics of the Knee



[image: image5.wmf]
Motion of the knee may be described in terms of rotations about or translations along some set of axes. "Rotation" refers to turning about an axis; "translation" is to slide along an axis, or more precisely, motion such that all lines in the body remain parallel to their original position. Conventionally, the axes are arranged to fit the "right hand rule" (that is, the X, Y and Z axes are arranged as the index finger, long finger and thumb of the right hand in the illustration). These translations and rotations are used to define motion. For example, motion along X is anteroposterior movement, along Y is medolateral inducement, and along Z is distraction or penetration of the femur. Tibia rotation about X is abduction or adduction, about Y is flexion-extension, and about Z is tibial rotation.

The ability to rotate or translate about a particular axis is called a "degree of freedom," but natural knee motions and the motions resulting from clinical laxity tests are not limited to one degree of freedom. For example, "knee flexion" during normal gait requires the knee to rotate about the medial/lateral (Y) axis, translate along and rotate about the anterior/posterior (X) axis and rotate about the proximal/distal (Z) axis (see illustration, pg. 12). A common way to think about how the knee flexes is to imagine the "axis of rotation" moving posteriorly. The location of this "axis" is often called the knee's "instant center".
	Axis
	Rotation
	Displacement
	Axis

	Anterior/

Posterior
	Abduction/Adduction
	Anterior/

Posterior
	X

	Medial/ Lateral
	Flexion/Extension
	Medial/ Lateral
	Y

	Proximal/

Distal
	Internal/External
	Proximal/

Distal
	Z


Coupled Motion and Constraint
The concepts of "constraint" and "coupled motion" are important to understanding normal movements of the knee. Constraint is absolutely essential to the normal function of the knee, to bracing, and to surgical reconstruction. As the knee moves from flexion to extension, the tibia rotates externally relative to the femur. The extension and external rotation are a coupled motion (see illustration, pg. 13). The knee is constrained by ligaments, tendons and by the joint capsule to a range of flexion of from 0° to about 150°. Similarly, the soft tissues constrain rotational motions to a range of 20°-30°. Movements into abduction or adduction are also limited. The  normal knee is constrained in how it can move. 

	Axis                     Rotation            Translation

	[image: image6.png]Penetration

External T l
; Distraction






	[image: image7.png]Extension

g
g
7
=
=






	[image: image8.png]Anterior

Adduction :.

A

v Posterior
. Abduction





The three axes of the knee. Around each axis there is a rotation and along each axis there is a translation to create a total of six degrees of freedom.



Knee motions are not independent of one another. For example, when the knee moves into extension, the rotation of the tibia on the femur is coupled to the movement of  ex-tension by knee construction and by the way in which ligaments, tendons and bones share the applied loads.

When testing joints one must take care not to provide inappropriate constraint. Joint laxity tests frequently focus on one motion. The constraint of the testing system on associated (coupled) motions may have a major effect on the movement being measured. For example, when performing the 20° anterior displacement test, if joint rotation is con-strained by the testing system, the tibia's anterior displacement will be significantly diminished.

Another example of coupled motion is the combined translation and rotation of the femur on the tibia as the knee flexes and extends. If the cruciate ligaments are approximately isometric (they do not stretch or shorten significantly as the knee bends) then the flexion angle determines the orientation of these ligaments. If one removes a constraint (for example, by cutting the anterior cruciate ligament) the coupling between the bones is changed. If one adds a constraint (a uni-axial brace) the coupling between the bones changes and new and different loads are applied to constraining structures. The components of the knee are linked by constraining and coupling structural components.

Flexion and the Ligaments
When flexion of the knee begins, the patellar tendon passes anteriorly from the tibial tubercle to the patella, relative to a line drawn perpendicular to the surface of the tibial plateau. As the knee flexes farther, the femur rolls posteriorly on the tibia under the control of the cruciate ligaments. The direction of the patellar tendon changes continuously from an anterior to a posterior orientation. When the quadriceps muscle is contracted, the forward movement of the tibia due to the anteriorly directed patellar tendon may be constrained by the anterior cruciate ligament (ACL). Similarly, the backward motion of the tibia that results from a posteriorly directed patellar tendon may be constrained by the posterior cruciate ligament (PCL). At the quadriceps neutral angle, however, the relationship of the tibia to the femur requires no cruciate restraints to the forward or backward motion of the tibia. The length of the cruciates remains approximately constant throughout the range of flexion. The term "isometric" has been used to describe these "unchanging dimensions" (see illustration, pg. 14). Although no ligament or tendon which bears a load can be literally isometric, when deformations are small, the approximation of isometry is observed.

The Cruciate Deficient Knee.

[image: image9.png]Anterior

Displacement
Force
N

Posterior
isplacement "

External
Rotation




Linked or joined motions are referred to as “coupled motions.”  An example of the coupled motion is the anterior displacement and internal rotation of the tibia that results from an anterior displacement force. A posterior displacement force also results in coupled motion, posterior displacement and external rotation in the tibia.

While the knees' mechanisms are multi-dimensional, a cruciate ligament deficiency can be detected in most knees by tibial sagittal plane displacements alone. A significant difference is usually seen in force/displacement behavior between normal and cruciate deficient knees. In a normal knee, anterior force results in anterior displacement and internal rotation of the tibia; posterior force results in posterior displacement and external rotation of the tibia. Since the patellar tendon is oriented posteriorly, a quadriceps contraction will not result in an anterior tibial shift in a normal knee positioned at 90° (although there might be a slight posterior shift). On the other hand, a contraction of the quadriceps with a PCL deficient knee at 90° will result in an anterior shift of the tibia of 2 mm or more.

The Knee Ligament ARTHROMETER® measures the relative motion between the patella and tibia. The tibial rotation motion is unconstrained. The abnormal tibial rotation that results from a ruptured ACL is reflected in sagittal plane measurements.

[image: image10.png]L

oL

FeL

Ac]

acL,

acL




The length of normal cruciate ligaments is approximately constant throughout the range of flexion,                                a state referred to as being “isometric”
Active Displacement. To generate sufficient torque to raise the foot, the patient contracts the quadriceps muscles. The quadriceps force acts on tibia through the patellar tendon. The patellar tendon force has two components: one compressing the joint surfaces together and one directed anteriorly. The anterior component acts to displace the tibia while the ACL acts as the primary restraint to its movement. When there is an ACL disruption, the anterior component may displace the tibia farther forward. For a more detailed discussion, see The Active Drawer, a booklet included with each ARTHROMETER®.

Quadriceps Neutral Angle. At the quadriceps neutral angle, the relationship of the tibia to the femur requires no cruciate restraints to the forward or backward movement of the tibia. The quadriceps neutral angle ranges from 55° to 90° and in a non-involved knee it is that angle where a quadriceps contraction does not produce either an anterior or posterior tibial shift. After a PCL tear, gravity can displace the tibia and will charge the normal orientation of the patellar tendon. When the patient contracts his quadriceps muscles during and active displacement test, the patellar tendon force will have an anterior component which pulls the tibia out of its posterior displaced position and up to the normal anatomical position. The Knee Ligament ARTHROMETER® is used to measure this posterior tibial sag in millimeters. This is done by measuring the active displacement on the involved knee at the same flexion angle that the non-involved knee was found to have no measurable displacement. Any resulting anterior tibial shift on the involved knee at the same degree of flexion is equal to the posterior tibial subluxation. For a more detailed discussion, see The Quadriceps Active Test, a booklet included with each ARTHROMETER®.

Getting Started


Arthrometer Accessories and Their Use
[image: image11.png]



The goniometer is used to measure and confirm flexion angles.
Goniometer. Two plastic rulers joined by a rivet pivot to indicate the angle of knee flexion. There are two principal methods of goniometer measurement. The first is to align the pivot point with the knee's joint line and to adjust one arm parallel to the femur and the other parallel to the tibia. The second method is to place the goniometer on top of the patella with the pivot aligned with the joint line; one arm rests on the anterior aspect of the tibia, the other on the anterior aspect of the femur. Although the first method is more precise, either will work. It is imperative, though, that the same method be used each time to ensure test-retest repeatability.


[image: image12.wmf]
The thigh support platform adjusts to accommodate patients of varying heights.
Thigh Support Platform. The adjustable platform is used to ensure  uniform right-left and test-retest angles of flexion in ACL displacement measurements. It is not used in PCL tests either at 90° or the quadriceps neutral angle. To adjust the height, pull on the platform's top casing; it will lock into one of three elevations. At its lowest elevation, the platform will raise the thigh 11 cm from the examination table surface, at its highest elevation it will raise the thigh 17.5 cm. 
Note: Early model KT1000™ thigh support platforms are not adjustable; to alter the elevation of one of these, place a board or book under the platform.
[image: image13.png]



Foot Support Platform. When placed under both feet of the patient, this accessory assures initial tibial rotation symmetry between both of the patient's limbs during ACL displacement testing. It is not used in PCL tests either at 90° or the quadriceps neutral angle. Position the patient's heels on the platform with lateral malleoli just proximal to the upright; this provides a restraint against the tendency for the patients entire leg to rotate externally.  When both feet are placed symmetrically in the support, the starting rotation angle of each limb will be the same. The foot support platform features a 10 cm ruler on either side; this corresponds to the " heel position " line on the Patient Evaluation Form; the position should be recorded at each test.  Follow up tests should replicate the angle of earlier exams to ensure validity of retest results.

Thigh Restraint Strap. A 7 cm Velcro strap is used to prevent excessive external hip rotation  during ACL displacement testing. It also may be used in PCL tests either at 90° or the quadriceps neutral angle. Place the strap around both the patient's legs at a level proximal to the thigh support platform. Tighten the strap until it removes much of the excessive external rotation when the feet are resting on the foot support platform (the patella should be pointing up). The thigh support strap is autoclaveable for intraoperative use. (Note: the KT1000™ is not autoclaveable; however, the KT1000™ can be protected by means of a sterile bag).
Setting Up for PCL Tests
1.  Relax the Patient. The patient should lie supine on a comfortable, firm examining table. The table should be long enough to support the patient's head and heels; the patient's head should be comfortable, supported with the cushion of the table or a pillow, and hands over his midsection (see  checklist, pg. 20). You may wish to use a modesty drape.
2.  Flex the Knee. Position knee to desired angle. Measure the angle of flexion using the goniometer and record on Patient Evaluation Form.
[image: image14.png]



An assistant at the patient’s side may help keep subjects leg in position and relax the patient.
3.  Position the ARTHROMETER®. Secure the Knee Ligament ARTHROMETER® on the anterior aspect of the tibia so that the patella reference pad rests only on the patella and remains totally proximal to the patellar tendon; apply the distal Velcro strap (without over tightening, which may cause discomfort and interfere with relaxation). Firmly secure the proximal Velcro strap.
4.  Adjust ARTHROMETER®. Confirm that the joint line arrow is at the level of the joint line and that pressure on the patella stabilizes the patella in the femoral trochlea. To adjust the ARTHROMETER® for the patella height, loosen the patella sensor adjustment knob and position the tibia and patella rails nearly parallel with the ARTHROMETER® case; tighten the knob.
5.  Restrain Leg. Maintain patient's leg in position either with the help of an assistant (who may sit by the patient's side to support leg and facilitate relaxation) or by sitting on patient's foot.


[image: image15.wmf]
Maintain hand pressure on patella throughout examination.
6.  Manually Oscillate the Calf. Grasp the back of the patient's proximal calf and gently but rapidly oscillate the tibia in an anterior and posterior direction several times; this will induce a relaxation response in the muscles. When the ARTHROMETER® dis-placement needle moves freely, the patient is relaxed.

7.  Stabilize the Patella. Place one hand on the thigh and the patella reference pad to stabilize the instrument. Push posteriorly on the patella reference pad until there is no further movement on the dis-placement dial. Maintain patellar stability throughout the exam.

8.  Position Tibia and Condition Tissue. Pull and then push on the force handle (this will produce  three anterior audio tones when pulling and the posterior tones when pushing); when the 20 lb. posterior tone sounds, release the force. Repeat. This posterior push and release maneuver is performed two or three times (until the displacement dial returns to the same position).

9.  Adjust Dial. While maintaining constant pressure on the patella reference pad, rotate the displacement dial to align the needle with zero.

Setting Up for ACL Tests
[image: image16.png]



Arthrometer® thigh support platform, foot support platform and thigh straps
1.  Relax the Patient. The patient should lie supine, with his hands at his side, on a comfortable, firm examining table. The table should be long enough to support the patient's head and heels; the patient's head should be comfortable, supported with the cushion of the table or a pillow, his hands over his midsection and his eyes looking to the ceiling. You may wish to use a modesty drape. (See checklist, pg. 20)

2.   Thigh Support Platform. Place thigh support platform under both legs at a level proximal to the popliteal space. This platform will keep the flexion angle of both knees symmetrical during testing. The thigh support platform is adjustable, raising the thigh an additional 6.5 cm (earlier production models are not adjustable; place a board or book under the platform to alter elevation).  Elevating the thigh support will increase the flexion angle. It is important to obtain an angle of flexion which well establishes the patella in the trochlea. Patellar mobility will produce test error.


[image: image17.wmf]
3.  Foot Support Platform. Place the foot support platform under both feet of the patient at a level distal to the lateral malleolus. The patient's heels should rest comfortably on the foot pad immediately adjacent to lateral upright ends. Note: Some patients must be tested with the heel resting on the foot pad at a distance in from the lateral uprights to prevent excessive internal rotation of the knee. In these cases, draw an imaginary line from the medial border of the foot down to the ruler on the foot pad. Read the actual distance that both heels are from the upright on the foot support's centimeter scale; record readings on the Patient Evaluation Form.

4.  Check Flexion Angle. Measure the angle of flexion using the goniometer; if necessary, adjust thigh support platform so that knee is between 20º and 35º of flexion. The knee must be flexed sufficiently for the patella to engage the femoral trochlea. Record the flexion angle on the Patient Evaluation Form.

[image: image18.png]



Align the Knee Ligament Arthrometer® with the medial joint line.
5.  Place Thigh Strap. Use the Velcro thigh strap if there is an excess of hip external rotation (either naturally or in the operating room when the patient is under general anesthesia). Place strap so it resists external rotation on both legs equally, then tighten until it removes much of the internal rotation torque on each knee when feet are in the foot support.

6.  Position the ARTHROMETER®. Secure the Knee Ligament ARTHROMETER® on the anterior aspect of the tibia so that the joint-line arrow aligns with the joint line of the knee; apply the distal Velcro strap (without over tightening, which may cause discomfort and interfere with relaxation). Rotate the instrument so that pressure on the patella pad stabilizes the patella. If the patella is unstable, elevate the thigh support further to flex the knee and bring the patella down into the femoral trochlea. Firmly secure the proximal Velcro strap.

7.  Adjust ARTHROMETER®. Confirm that the joint line arrow is at the level of the joint line and that pressure on the patella stabilizes the patella in the femoral trochlea. To adjust the ARTHROMETER® for patella height, loosen the patella sensor adjustment knob and position the tibia and patella rails nearly parallel with one another; tighten the knob. 

8.  Manually Oscillate the Calf. Grasp the back of the patient's proximal calf and gently but rapidly oscillate the tibia in an anterior and posterior cycle several times; this will induce a relaxation response in the muscles. When the ARTHROMETER® displacement needle moves freely, the patient is relaxed.

[image: image19.png]



Secure the Arthrometer® to the subject’s leg using the Velcro straps.
 9.  Stabilize the Patella. Place one hand on the thigh and the patella reference pad to stabilize the instrument. Push posteriorly on the patella reference pad until there is no further movement on the displacement dial. Maintain this pressure throughout the examination.
10.  Position Tibia and Condition Tissue. Pull and then push on the force handle (this will produce three anterior audio tones when pulling and posterior tones when pushing); when the posterior tone sounds, release the force. Repeat. This posterior push and release maneuver is per-formed two or three times (until the displacement dial returns to the same position).

[image: image20.png]



To zero displacement dial, rotate the bezel until needle is aligned with “0”
11.  Adjust Dial. While maintaining constant pressure on the patella reference pad, rotate the displacement dial to align the needle with zero.

[image: image21.png]



While maintaining patella pressure with one hand, pull and push on the force handle to position tibia and condition tissues.
The Testing Reference Position

Once the dial has been "zeroed," establish the testing reference position.  As you maintain posterior pressure on the patella reference pad with one hand, push on the force handle with your other hand until you hear the 20 lb. force tone; release the handle.  Rotate the dial until the zero aligns with the needle as you maintain pressure on the patella pad. Repeat until you return to the same position three time.
The testing reference position is the zero position from which all other measurements, posterior and anterior, will be made.  Failure to establish a repeatable testing reference testing position may indicate a lack of patient cooperation (tense thighs, legs held together or rigid, tight patellar tendon during manual palpitation).

Patient Relaxation Checklist
· Oscillate Calf. Watch the dial: are readings higher than when making passive force measurements?

· Drop Thumb to Patellar Tendon. Is there tension in tendon?

· Obtain Zero Reading from Posterior Push. Does dial return to a more anterior reading after repeated push?

· Ask Patient if He/She is comfortable. Does testing cycle hurt?

· Is Patient Watching Test? Eyes should be at ceiling.

· Patient's Hands. Are they over the midsection?

· Feel Patella With Second (Index) Finger Under Patella Paddle During Passive

· Testing.  Does patella  slide during testing? Elevate thigh support.

· Feel Quads During Passive Test. Is there tension in quads?

· Try Using Thigh Strap.
· Try Using Modesty Drape.

Other Considerations
· Excessive Soft Tissue at Proximal Calf? Does tissue impinge tibial paddle during passive anterior testing? Move strap distally.

· Does Patient Wince with Posterior Force?
· Is Patient Wearing Shoes?

· Is Patient Wearing Long Trousers?

· Is Quadriceps Active Test at 20°-30° Greater than 30 lb. Test? Patient is guarding.

· Does Maximum Dial Excursion During Oscillation of Calf Exceed 20 lb. Passive Tests? Patient is guarding.

Setting Up the X-Y plotter (KT2000™ Only)
[image: image22.png]



The KT2000™ X-Y plotter
A - Power

B - Off-Chart-Pen Switch

C - "X" Adjustment Knob

D - "Y" Adjustment Knob

E - Pen

1. Connect to ARTHROMETER®. Connect coiled cord to plotter's five-pin, circular connector and to the rear jack of the ARTHROMETER®.

2. Insert Patient Evaluation Form. 

a) Move plotter's power switch (the red switch on the lower left corner on top of the plotter) to the "1" position.

b) Place a KT2000™ Patient Evaluation Form on the plotting surface, grid side up with the lower edge of the form against the surface edge.

c) Move the Off-Chart-Pen switch to the "Chart" position. Make sure paper is on the plotting surface before switch is moved to "Chart;" retention force of the surface will be reduced if the plotter is on "Chart" without paper on this surface.

3. Drop the Plotter Pen. Depress the switch on either side of the Knee Ligament ARTHROMETER® force handle; the pen will remain in position as long as the switch is held.

4. Zero the Plotter. Push on the force handle until you hear the 20 lb. tone; release the force handle. The pen should return to the 0 force line on the graph paper. (If it does not, move the plotter pen up or down on the paper using the vertical adjustment knob on the upper right of the plotter until the pen drops directly onto the 0 force line). After the initial posterior push, you will have only 40 seconds to zero the plotter (the ARTHROMETER® turns itself off automatically). If necessary, give the unit another push to re-start it.

5. Place Pen in "Test" Position. Using the horizontal adjustment knob, place the pen in the desired position to begin the force/displacement curve. Each graph form has ample room for testing and comparing left and right knees. Up to three curves on each knee (six total readings) may be recorded on a single page.

Setting Up the X-Y Plotter (Recent Model KT2000™)


A - Power Pen Up/Down Lever

B - Paper Clamp

C - Pen

D - "X" Adjustment Knob

E - "Y" Adjustment Knob

1. Connect to ARTHROMETER®. Connect coiled cord to plotter's five-pin, circular connector and to the rear jack of the ARTHROMETER®.

2. Insert Patient Evaluation Form.

a) Move plotter's power switch (the red switch on the lower left corner on top of the plotter) to the "load" position. This moves the pen-holder to the upper-right corner and releases the right-hand paper clamp. 

b) Release the left-side paper clamp by depressing the tabs. Remove the stub or cardboard backing remaining from the previous pad.

c) Lay the new pad on the surface and slide it well under the right-side paper clamp. Depress the left-side paper clamp and slide the new pad under it from right to left. The pad should line up against the left paper stop pin (under the left-side clamp) and the bottom edge of the paper stop. Release the left-side clamp. Smooth pad against surface by moving your hand from left to right over the pad.

3. Drop the Plotter Pen. Move the plotter's operating switch to the Pen-Down position.

4. Zero the Plotter. Push on the force handle until you hear the 20 lb. tone; release handle. The pen should return to the 0 force line on the graph paper. (If it does not, move the plotter pen up or down on the paper using the vertical adjustment knob on the upper right of the plotter until the pen drops directly onto the 0 force line). After the initial posterior push, you will have only 40 seconds to zero the plotter (the ARTHROMETER® turns itself off automatically). If necessary, give the unit another push to re-start it.

5. Place Pen in "Test" Position. Using the horizontal adjustment knob, place the pen in the desired position to begin the force/displacement curve. Each graph form has ample room for testing and comparing left and right knees. Up to three curves on each knee (six total readings) may be recorded on a single page.

Testing Technique
To achieve reproducible results, users of the Knee Ligament ARTHROMETER® must develop a reliable and reproducible testing technique. The so-called "pull-push" cycle works effectively for users both the KT1000™ and KT2000™. An alternative technique, the "push-pull" cycle, has been developed by users of the KT2000™. With practice and study, either method will yield accurate and reproducible results.

Pull-Push testing Cycle (KT1000™)
1. Set up Patient for PCL or ACL Tests.

2. Establish the Testing Reference Position.

3. Maintain Pressure on the Patella Pad throughout Cycle.

4. Pull: smoothly and slowly through the force handle (with the pull in line with the handle shaft); tones will sound at 15 lb., 20 lb. and 30 lb. of force. Do not exceed 30 lb. of force.

[image: image23.png]P L ——

s




X-Y Plot of pull-push cycle. This graph indicates anterior displacement of 9 mm at 30 lb.,; 8 mm at 20 lb.. and 7 mm at 15 lb.. and posterior displacement of 2 mm at 20 lb.. 

Note: Excessive application of force beyond the 30 lb. anterior and 20 lb. posterior tones will result in patient guarding due to discomfort at contact points on the tibia and possible pain from excessive force.

5. Push: on the force handle until the 20 lb. posterior tone sounds. Do not exceed 20 lb.. of force.

6. Release Force Handle. The dial should return to "0", ± 0.5 mm. This indicates the orientation of  the instrument is unchanged and the patient has remained relaxed.

7. Record Readings. Record displacement readings at each tone on the Patient Evaluation Form.

8. Repeat (If Necessary). If the dial does not return to within 0.5 mm of zero, do not use measurements; re-establish testing reference position and repeat cycle.

Pull-Push Testing Cycle (KT2000™)
Note: The KT2000™ has a specialized battery saving feature which automatically shuts off the ARTHROMETER® battery power after 40 seconds of inactivity. To re-start the ARTHROMETER® push on the force handle until you hear the posterior force tone; then continue with your tests.
1. Set up Patient for ACL or PCL Tests.

2. Establish the Testing Reference Position.

3. Maintain Pressure on the Patella Pad throughout Cycle.

4. Zero Plotter. Adjust X-Y plotter to zero.

5. Pull: while depressing the pen button (in the force handle), pull smoothly and slowly through the force handle (with the pull in line with the handle shaft); tones will sound at 15 lb.., 20 lb.. and 30 lb.. of force. Do not exceed 30 lb.. of force.

6. Push: on the force handle until the 20 lb. posterior tone sounds. Do not exceed 20 lb.. of force.


[image: image24.wmf]
X-Y plot of push-pull cycle. This graph indicates anterior displacement of 9 mm at 30 lb..; 8 mm at 20 lb.. and 7 mm at 15 lb. and posterior displacement of 2 mm at 20 lb.. 

Note: Excessive application of force beyond the 30 lb. anterior and 20 lb. posterior tones will result in patient guarding due to discomfort at contact points on the tibia and possible pain from excessive force.

7. Release Force Handle. While maintaining pressure on the pen button, release posterior force; the dial should return to "0", ± 0.5 mm. Remove pressure from the pen button.

8. Repeat (If Necessary). If the pen does not return to within 0.5 mm of zero, do not use 

measurements; re-establish testing reference position and repeat cycle.

9. Measure. Place the clear, plastic reticular scale on the plotter graph. If the scale is missing, count vertical lines (from the zero reference point); each line equals 1 mm. Displacement to the left of zero (or the starting reference position) indicates posterior displacement; displacement to the right indicates anterior displacement.

Push-Pull Testing Cycle (KT2000™)
[image: image25.png]weores Copwen
ooz

N

LA R R

e ]
T234567800M02
N ]

N o = =
IR R R EERETETIEY





KT2000™ reticular scale aligns at 15, 20 and 30 lb.. anterior, 20 lb.. posterior and displays millimeters of displacement. 

Note: The KT2000™ has a specialized battery saving feature which automatically shuts off the ARTHROMETER® battery power after 40 seconds of inactivity. To re-start the 

ARTHROMETER® push on the force handle until you hear the posterior force tone; then 

continue with your tests.

1. Set up Patient for ACL or PCL Tests. (See pages 16, 17, 20)
2. Establish the Testing Reference Position. (See page 19)

3. Maintain Pressure on the Patella Pad throughout Cycle.
4. Zero Plotter. Adjust X-Y plotter to zero.

5. Push. Push on the force tone handle until the posterior 20 lb. tone sounds. Depress button on force handle. Do not exceed 20 lb.. of force.

6. Pull. Pull forward on the force tone handle until the 30 lb. anterior tone sounds; release button to lift the plotter pen. Do not exceed 30 lb.. of force.

7. Push. Push posteriorly again until the 20 lb. posterior tone is heard; release the force handle while continuing to stabilize patella. The instrument should return to zero ± 0.5 mm; this confirms that the instrument testing reference position has not changed.

Note: Excessive application of force beyond the 30 lb. anterior and 20 lb. posterior tones will result in patient guarding due to discomfort at contact points on the tibia and possible pain from excessive force.

8. Measure. Place the clear, plastic reticular scale on the plotter graph. If the scale is missing, count vertical lines (from the zero reference point); each line equals 1 mm. Displacement to the left of zero (or the starting reference position) indicates posterior displacement; displacement to the right indicates anterior displacement.

Testing Cruciate Ligaments

90 Degree Quadriceps Active Test (Visual)

Purpose: To test for posterior cruciate ligament disruption. 

1. Position Patient. Have patient lie supine on examining table. Have patient flex knee, keeping your eyes at the level of the flexed knee; rest one elbow on the table, using your hand to support the patient's thigh and to confirm that thigh muscles are relaxed. (See pages 16 and 20. )

2. Apply Force. Using the other hand, apply a brief posterior force to the tibia and then lightly restrain the patient's foot. Ask patient, "gently slide your foot down the table;" your hand on the foot will prevent the leg from sliding. Anterior tibial displacement resulting from the quadriceps contraction indicates that the PCL is disrupted.  

Quadriceps Contraction

[image: image26.png]



Performing 90 degree Quadriceps test.  The tibia of a knee with an injured PCL will  move anteriorly. This is because the patellar                                               tendon is oriented anterior as it passes  from the posterior subluxed tibia to the patella

3. Quantify 90° Quadriceps Active Findings: This is a more sensitive test to screen for PCL disruption. 

a) Set up Patient for PCL Testing. 

b) Establish Testing Reference Position. 

c) Test the Non-Involved Knee. Ask patient to "Gently try to slide your foot down the table;" note displacement on the dial. Apply a momentary 20 lb. posterior push on the force handle. The dial needle should return to zero ± 0.5 mm. 

d) Test the Involved Knee. Repeat sequence (a-c above) on Involved Knee. 

e) Interpreting Findings. Anterior movement of the tibia indicates an injury to the posterior cruciate ligament; posterior movement from 0.5 mm to 2 mm of the tibia indicates a normal knee. 

4. Follow-up Testing. If the tibia moves anterior during quadriceps active testing, perform further PCL tests at the quadriceps neutral angle. If quad active tests do not indicate a PCL, perform ACL tests. 

Tests at the Quadriceps Neutral Angle

Purpose: To evaluate the PCL injured knee. 


[image: image27.wmf]

   TIBIAL

SAG TEST

ACL

TESTING

INTACT

PCL

ACL

TESTING

ABSENT

PCL

Anterior Displacement

Active

Sick Knee



A/P Displacement

Passive





30







20







15







Compliance

Index





Find QNA

Well Knee

at = 70º



Fail

+ Value





A/P Displacement 30º

(Passive)

L                               R



15







20







30







Compliance

Index





Manual

Displacement

L                   R



A/P Displacement 30º

(Active)

L                                R



Pass

- Value





Visual Screen

for PCL at 90º





Active Screen

for PCL at 90º






1. Set up Patient for PCL Testing. 

2. Establish Testing Reference Position. 

3. Establish the Quadriceps Neutral Angle on the Non-Involved Knee. Ask patient to "Gently try to slide your foot down the table" (while restraining the foot). Both the position of the foot and the flexion angle of the knee should remain constant. If there is more than 1 mm of deflection on the dial, adjust the knee flexion angle and repeat until the flexion angle is identified that has no more than 1 mm of tibial motion with the quadriceps active test. This is the quadriceps neutral angle (in studies, this angle has ranged from 55° to 85°). Measure the angle of the knee with the goniometer and record the quadriceps neutral angle on the Patient Evaluation Form.  

[image: image28.png]INJURED KNEE

Messured Laxiy Carrected Lasty
Quitris Actve
poston (GAPY

epositon)

HHHP




4. Perform the Passive Displacement Test on the Non-Involved Knee. Perform the standard pull-push testing cycle. Record testing angle, 20 lb. anterior and 20 lb. posterior displacements on the Patient Evaluation Form. 

5. Perform the Quadriceps Active Test on the Non-Involved Knee. Zero the ARTHROMETER® and have the patient slide the restrained foot forward; record the result on Patient Evaluation Form. The dial should read zero (at QNA).

Displacement measurements at the quadriceps neutral angle. The examples might be typical of an isolated PCL injury (i.e. the ACL is intact.)
	
	Injured 

Knee Measured Laxity
	Injured Knee

corrected laxity
	Normal

 Knee

measured laxity
	Injured-Normal

I-N laxity

	20 lb. anterior
	12.0
	2.5
	2.0
	0.5

	20 lb. posterior
	3.5
	13.0
	1.0
	12.0

	Quadriceps active displacement
	9.5
	
	0
	


6. Perform the Passive Displacement Test of the Involved Knee. Perform the standard pull-push testing cycle. Record testing angle, 20 lb. anterior and 20 lb. posterior displacements on the Patient Evaluation Form. 

7. Perform Quadriceps Active Test on the Involved Knee. Zero the ARTHROMETER® and have the patient slide the restrained foot forward; record result on Patient Evaluation Form. 

8. Adjust Measurements. Normalize displacements to the quadriceps active position. The corrected anterior displacement is displacement anterior to the quadriceps active position (measured anterior displacement minus the quadriceps active displacement). The corrected posterior displacement is displacement posterior to the quadriceps active position (measured posterior displacement plus the quadriceps active displacement). 

9. Interpret Data. A difference of 2 mm or greater usually indicates pathology. 

10. Follow-up Testing. Proceed to ACL tests. Note: Only the compliance index passive

displacement test may be used in the case of the PCL disrupted knee. 

For further information on instrumented PCL testing, refer to the video program, Instrumented Drawer Testing, or to the booklets, The Quadriceps Active Test, and, The Active Drawer Test. All are included with the KT1000™ and KT2000™; additional copies are available directly from MEDmetric® Corporation. 

[image: image29.png]POSTERICR DISPLACEMENT AT QUAD NEUTRAL ANGLE [-707)

POSTERIOR

Passie Depacemert Tees (i rfced)
(Ficin= 7D, “vhen Quadices s Diplcemart NeD)

Force

]

EDS

Qusdcaps Ace Dplcement (Tt 535 Tos)
(Pt sides restained ot sing e g
B igh strap for lteral suppar)

Canected 2 b, Posiarion Diplacement
Akt ACTVE Dapiscement o PASSIVE Potsr Dispacoment

Carected 20 b. Antarr Dislcament

Sublract ACTIVE displacement from PASSVE Anterior Displacemant





[image: image30.png]



Quadriceps neutral angle test  measurements and adjustments as recorded on the Patient Evaluation Form (above) and the KT2000™ X-Y plotter. Note that in this case the quad neutral angle is at 70° of flexion; studies have indicated quadriceps neutral angles ranging from 55° to 85°.
Passive Displacement Test

[image: image31.wmf]ACL TESTS

Involved Knee







Manual Max Test

Non-Involved Knee





Involved Knee







Quadriceps Active Test

Non-Invoved Knee





Determine

Compliance Index





Involved Knee







Passive Displacement Test

Non-Involved Knee






Purpose: To test for anterior cruciate ligament deficiencies. 
1. Set up Patient for ACL Testing. 

2. Establish the Testing Reference Position. 

3. Test the Non-Involved Knee. Mark the medial joint line; perform the standard pull-push test.  Record measurements at 15 lb., 20 lb. and 30 lb. of anterior force on the Patient Evaluation Form. 

4. Test the Involved Knee. Support the involved knee at the non-involved knee test flexion angle; establish the testing reference position for 30º. Perform standard pull-push test; record 15 lb., 20 lb. and 30 lb. anterior displacements. 

5. Interpreting Data. In most individuals, the measured 

difference between knees is usually less than 3 mm.
[image: image32.png]ANTERIOR DISPLAGEMENT MEASUREWENT AT 20" TO 35"

ANTERIOR POSTERIOR

b,

pouaren g = D0  psiposin = O o) [ 0.

I

[ T —— e W
@i st trugh KTIQE0™ ) e 5





[image: image33.png]



The results of 30º passive displacement tests as recorded on the Patient Evaluation Form (above) and the X-Y plotter grid (below). The graph on the left indicates the non-involved knee; the one on the right the involved knee.
6. Possible Inconsistencies. Force vectors, rotation of the ARTHROMETER® or parallax may create measurement inconsistencies. If measurements are inconsistent on the 30° tests, check the following: 

· Transient Muscle Guarding (Solution: Ensure patient is relaxed; manually oscillate calf to further relax muscles) 

· Variation in the Degree of Initial Tibial Rotation (Solution: Ensure test is started from same position of tibial rotation through use of the thigh strap) 

· Variation in Patella Pad Pressure (Solution: Check starting pad pressure; establish patella in trochlea) 

· Failure to "Zero" between Measurements. (Solution: Give a momentary posterior push to reset the tibia at the test reference position). 

· Inconsistent Knee Flexion Angle. (Solution: Compare angle on goniometer with previously recorded angle). 

· Skin "Bumping" Error. This may be a problem in an obese patient or one with mobile skin and soft tissue around the calf. (Solution: Slide proximal Velcro strap more distal - toward the foot). 

· General Anesthesia Effects in the O.R. (Solution: Lower foot of the O.R. table to bring flexion angle closer to 35° for both knees; this will guide the patella into the trochlear groove).  Note: See "Sources of Measurement Error" 

7. Follow-up Testing. Compute the Compliance Index. 

The Compliance Index
Purpose: To quantify the clinical feeling of a "soft endpoint" observed after an anterior cruciate ligament disruption. This is done using data gathered from 30° passive displacement tests. Compare both knees using this test at every follow-up evaluation. 

Subtract 15 lb. readings from 30 lb. readings. 

The resulting measurements constitute the compliance index, the slope of the

force-versus-displacement curve of the patient's knee between 15 lb. and 30 lb. of applied anterior force.

Proceed to the Quadriceps Active Test

[image: image34.png]ANTERIOR DISPLACEMENT MEASUREMENT AT 20° TO 35 ANTERIOR

e r— R
(Examines applies force through KT1000™ handse) e & (G O [
o] 9 g«

vt 30 s weirotn- @ en [ v 7513 |
7

([ compame e oy b 156 | 7,87 5





In determining the compliance index use data gathered during passive displacement tests
Quadriceps Active Test at 30°



[image: image35.wmf]
Purpose: To Measure tibial displacement with a quadriceps muscle contraction 

1. Set up Patient for ACL Testing. 

2. Establish the Testing Reference Position 

3. Test the Non-Involved Knee. Ask patient to "Gently try to raise your foot off the table." Just as heel lifts, note displacement on the dial and record on the Patient Evaluation Form. Apply a momentary 20 lb. posterior push on the force handle. The dial needle should return to zero ± 0.5 mm This test may also be performed by a restraining hand on the ankle. 

4. Test the Involved Knee. Repeat the sequence in Step 3 on the involved knee. 

5. Interpret Results. A right/left difference of 3 mm probably indicates a pathology. 

6. Follow-up Testing. Conduct manual maximum displacement test.

[image: image36.png]| e e Dpcemen Tess [ w] o
[ p——— el7





Quadriceps active test at 30° results on the Patient Evaluation Form (above) and the KT2000™ X-Y plotter (below) 

[image: image37.png]



Manual Maximum Displacement Test
[image: image38.png]



Purpose: To obtain maximum displacement and quantify endpoints. 

1. Set up Patient for ACL Testing. 

2. Establish the Testing Reference Position. 

3. Test Non-Involved Knee. Use one hand to stabilize patella; use the other to apply direct manual force to the back of the subject's calf, proximal to the instrument's Velcro straps. Read and record displacement from the dial. 

4. Zero the ARTHROMETER®. Apply posterior force through the handle and release. The dial needle should return to zero ± 0.5 mm. 

5.  Test the Involved Knee. Repeat the sequence in Step 2 on the involved knee. 

6. Interpreting Results. The Manual Maximum Displacement Test has been the most sensitive of the various tests. A right/left displacement measurement difference of 3 mm or more is indicative of an ACL disruption. 

[image: image39.png]Manucl Maximum Displecement Testa
(High force appled manualy to proimal calf





Manual maximum displacement testing as recorded on the Patient Evaluation Form (above) and KT2000™ X-Y plotter form (below). 

[image: image40.png]



Data Interpretation


A right/left difference of 3 mm in any measurement probably indicates a pathologic condition. If the 90° quadriceps active test is negative, indicating that the posterior cruciate ligament is intact, the  difference is probably due to an anterior cruciate ligament injury. Pathologic displacement due to ACL injury is best measured with the 30° tests. If the PCL is disrupted, the quadriceps neutral angle can be used to separate anterior displacement and posterior displacement. If the PCL is disrupted, the 30° tests can be used to measure total anterior-posterior displacement at that angle. However, one cannot discern how much anterior displacement is constrained by the ACL and posterior displacement is constrained by the PCL. The anterior compliance index (displacement between a 15 lb. and 30 lb. pull) tested at 30° can be used to diagnose an ACL disruption in the PCL-intact and PCL-disrupted knee. An anterior compliance index right/left difference of 3 mm is indicative of an ACL disruption. 
Sources of Measurement Error
Inconsistencies in measurements can occur from a number of sources: Force vector, rotation of the ARTHROMETER® or parallax. Among the reported sources of measurement error (and ways to correct them) are as follows. Sources of inconsistency unique to ACL testing procedures are above in the section about passive displacement testing. 

1. Inadequate stabilization of the patella. 

Stabilize patella by rotating limb to orient the patella anterior (superior pole of the patella pointing up). This may require a thigh strap. Support limb in sufficient knee flexion to engage patella in femoral trochlea. This may require elevating the thigh support. Apply constant pressure on the patella to seat it securely in the femoral trochlea. 

2. Inadequate instrument stabilization. 

 Prevent the instrument from rotating during the test by using the patella stabilization hand to steady the instrument about the tibial axis. Keep the hand in contact with the thigh. 

3. Incorrect placement of the ARTHROMETER®. 

 Confirm that the instrument joint line marker is on the joint line. 

4. Patella or tibia sensor rail motion limited by ARTHROMETER® case. 

 Solution: Adjust the height of the patella sensor pad. 

5. Inadequate patient relaxation. 

 Thigh muscles must be relaxed. 

6. Incorrect direction of force application. 

Force application should be in line with the force handle. 

7. Soft tissue displacement of the tibia sensor pad. 

On rare occasions in an obese patient, while an anterior force is being applied the proximal Velcro strap will cause the subcutaneous fat to displace anteriorly and lift up the tibial sensor pad. Applying the proximal strap more distal may solve the problem. 

8. Instrument malfunction. 

Check calibration. (See Maintenance and Troubleshooting). 

References
Daniel DM, Malcom LL.,, Losse G, Stone ML, Sachs R, Burks R. Instrumented measurement of anterior laxity of the knee. J Bone Joint Surg 1985; 67A:720-726. 

Daniel DM, Stone ML, Instrumented measurement of knee motion. In Daniel DM, Akesen WH, O’Connor JJ, eds. Structure, Function, Injury and Repair. New York: Raven Press, 1990; 420-426.

Daniel DM, Sone ML, Barnett P, Sachs R. The use of the quadriceps active test to diagnose posterior cruciate ligament disruption and measure posterior laxity of the knee. J Bone Joint Surg 1985; 70A:386-391. 

Daniel DM, Stone ML, Sachs, R, Malcom LL. Instrumented measurement of anterior knee laxity in patients with acute anterior cruciate ligament disruption. Am J Sports Med 1985; 13:401-407. 

Daniel, DM, Stone, ML, Rannger C. Instrumented measurement of anterior-posterior knee motion. In Aichroth PM, Cannon WD, Jr., eds. Knee Surgery - Current Practice. London: Martin Dunitz Ltd., 1992: 191-205.

Documenting Knee Stability in the O.R.
Documenting Knee Stability in the O.R. he Knee Ligament ARTHROMETER® can be used to document the course of knee stability, the natural   history of ACL tears and repair. The instrument is being used to verify chronic ACL tears. Under anesthesia, KT ARTHROMETERs® can document the pre-op stability and stability after reconstruction. Be sure that knee flexion and foot rotation angles are recorded and reproduced for each subsequent test. 

Using the KT1000™/KT2000™ ARTHROMETER® in the operating room, you can answer the following questions: 

1.  How much injured-to-normal difference was there pre-reconstruction? 

2.  How much was the displacement of the injured knee decreased by reconstruction? 

3.  Has normal anterior stability been established? 

If you plan to use your Knee Ligament ARTHROMETER® in the O.R., protect it from prep solution stains. Keep it out of the operative field and remember that it is an electronic device and cannot be sterilized. A specialized intra-operative Knee Ligament Arthrometer®, the KT1000/S™, is available. The KT1000/S™ is a mechanical device which may be sterilized using any common method.  Contact MEDmetric® at (800) 995-6066 for details. 

Note: Post-operative testing and management of the patient in rehabilitation should be undertaken under the direction of the operating surgeon.

Maintenance and Troubleshooting


The Knee Ligament ARTHROMETER® has been factory calibrated and will not require any further adjustment. Barring component failure, the force level audio indicators and the displacement gauge will remain within calibration throughout the instrument's lifetime. 

Your Knee Ligament ARTHROMETER® does not need periodic servicing beyond the replacement of standard 9-volt batteries found in the rear of the ARTHROMETER®. A weakening of the “beep” tone is evident when the batteries are near discharge. Weakened batteries affect KT2000™ plotter/computer output current. For this reason, force tones will cease altogether in the newest Models KT2000™ (Serial Number 327 and higher) when batteries are weak. Replace batteries with equivalent alkaline 9-volt batteries such  as Eveready, Duracel or Mallory. As the batteries weaken with use, you need not worry about the accuracy of your measurements since the Knee Ligament ARTHROMETER® automatically maintains calibration, even with a nearly dead battery.  To conserve battery life, you will note that there is no separate on/off switch. As you apply an anterior or posterior force through the force handle, the battery is automatically switched on. When the instrument is not in use, the battery is automatically in the off mode. For this reason, when you store your ARTHROMETER®, be sure that the force handle is not depressed. 

The Knee Ligament ARTHROMETER® will not need maintenance except as a con-sequence of component failure or damage. Note the case is factory sealed and cannot be opened in the field without causing internal damage. There are no user-serviceable parts or adjustments in the KT1000™ or KT2000™. Should it need repair, the instrument should be returned, along with its carrying case and accessories, to: 

MEDmetric® Corporation 

7542 Trade Street 

San Diego, California 92122-2412.

It is in both our interests that your unit provide trouble-free operation; we will promptly perform any necessary repairs and return the unit to you (see warranty information for further details).  

One final tip to extend the life of your instrument: Always keep and use the  ARTHROMETER® in a clean, dust-free environment, away from all fluids (such as prep solutions, alkaline and cleansing compounds). Also, keep your KT out of direct sunlight; extreme heat may warp the case. 
Force Tone Checks

Your Knee Ligament ARTHROMETER® should be tested for accuracy from time to time. The force tone tests will help you determine if your ARTHROMETER® audio tones are synchronized to the forces applied to the Arthrometer® handle.  You will need to place your Arthrometer® on a scale or balance.  Zero the scale o compensate for the weight of the instrument. 
Posterior Force Tones
If using a balance: 

1. Set the weights at 15 lb. greater than the weight of the KT. 

2. Push on the handle until the first tone is heard. 

3. The scale should indicate 15 (± 1.5) lb. [6.8 (± .68) kg].* 

4. Set the weights at 20 lb. greater than the weight of the KT. 

5. Push the handle until the second force tone is heard. 

6. The scale should indicate 20 (± 2.0) lb. [9.1 (± .91) kg]. 

If using a scale: 

1. Zero the scale to adjust for the weight of the KT. 

2. Push on the handle until the initial tone is heard. 

3. The scale should indicate 15 (± 1.5) lb. [6.8 (± .68) kg).* 
4. Continue to push on the handle until the second force tone is heard. 

5. The scale should indicate 20 (± 2.0) lb. [9.1 (± .91) kg]. 

Anterior Force Tones
A fish scale such as the G96 Mariner Model No. 2092 (available from sporting goods outlets) is suitable for this measurement. 

1. Hold the instrument on a table or similar flat surface. 

2. Hook the fish scale under the handle at the shaft. Restrain the instrument by placing one hand on top of ARTHROMETER® to provide counter-force. 

3. Apply upward force until the first force tone is heard. 

4. The scale should indicate 15 (± 1.5) lb. [6.8 (± .68) kg]. 

5. Continue to apply anterior force until the second tone is heard. 

6. The scale should indicate 20 (± 2.0) lb.. [9.1 (± .91) kg]. 

7. Continue to apply anterior force until the third tone is heard. 

8. The scale should indicate 30 (± 3.0) lb.. [13.6 (± 1.4) kg]. 
* The KT2000™ has no 15 lb. posterior tone.

Checking Accuracy







Fold Back Straps

[image: image41.png]



A. Goniometer 

included with

Arthrometer

B. Set pad flat on

level surface

[image: image42.png]ZERO DISPLACEMENT GAUGE
BY ROTATING BEZEL

ALIGN RULED EDGE
WITH JOINT LINE
ARROW; ADJUST
TO Sem




[image: image43.png]



DISPLACEMENT GAUGE SHOULD

INDICATE 10 mm POSTERIOR DISPLACEMENT (± 0.5 mm)
[image: image44.png]MANUALLY RAISE TIBIAL ARM TO 6em
MEASURED AT THE JOINT LINE ARROW





Checking for Linearity
[image: image45.png]'WITH STRAPS FOLDED
BACK, PLACE KT FLAT
ON & LEVEL SURFACE.
AND HOLD TIBIAL
AND PATELLAR

ARMS TOGETHER





[image: image46.png]



ZERO DIAL BY ROTATING BEZEL

[image: image47.png]LIFT TIBIAL AND PATELLAR ARMSTOGETHER AND OBSERVE DIAL
INDICATOR; THERE SHOULD BE
NO MORE THAN 05mm OF MOTIO]

HOTE: Ralative sncton
betveent il

and patllarsemer

pads rill podce

il mormanant




Procedures for the X-Y Plotter


Note: The graph paper provided by MEDmetric® Corporation is recommended as an accuracy standard. If you use graph paper from some other source (including photocopies of MEDmetric® Patient Evaluation Forms), the printing scale accuracy is questionable. The clear, plastic reticular scale provided with the KT2000™ may be used to gauge force and displacement accuracy. Remember to activate the ARTHROMETER® with a push prior to testing.  

Force Test: Checking Plotter Accuracy.
1. Connect ARTHROMETER® and X-Y plotter; place the plotter's pen on the “0” point. 

2. Momentarily press the button on the ARTHROMETER® force handle as you pull; release button.  Depress and release the button when the 15, 20 and 30 lb. anterior beeps sound. 

3. Momentarily press the button on the force handle as you  push; depress and release button when the 20 lb. posterior beep sounds. 

The graph paper should have pen marks at the coordinates corresponding to the anterior and posterior beeps. Each graphic mark should be ±10% of applied force (i.e., ±2 lb.. posterior and ±1.5 lb.., ±2 lb.. and ±3 lb.. posterior). 

Displacement Test: Comparing the X-Y Plotter and ARTHROMETER® Dial.

1. Adjust pen.  “Zero” pen horizontally on page. 

2. Lift on the arm connecting the patella pad to the ARTHROMETER®.  When the dial reaches 5 mm, 10 mm and 15 mm of displacement, momentarily press (then release) the button on the force handle. 

3. Lift the shorter arm (to the tibial pad).  When the dial reaches 5 mm, 10 mm and 15 mm of displacement, press and release the button on the force handle. 

The graph paper should have pen marks at the coordinates corresponding to 5 mm, 10 mm and 15 mm from the “0” position. Each graphic marks should be ±0.5 mm of applied displacement. 

Cleaning the Electrostatic Paper Holder.  Use lint free tissue (such as Kim wipes) and an anti-static, conductive cleaner (such as that made by Chemtronics). Moisten surface and wipe.  

Fuse Replacement.  Find the knob which corresponds to power (110v/220v); turn it

counterclockwise.  The knob doubles as a cover for the fuse.  Replace with a ULG/6.3 f x 32 mm fuse (outside the U.S., specifications may vary; consult your X-Y plotter user's guide for complete information) . 

Pen Replacement.  The pen slides in and out of its holder; insert the cartridge with the point down and slide in until you hear  a gentle “click.”. 

Power. Always turn off the power to the X-Y plotter when it is not in use (prolonged running may hamper effectiveness of electrostatic paper holder).

Limited Warranty

This warranty is limited to the original purchaser. 

MEDmetric® products are warranted by MEDmetric® Corporation to the original purchaser to be free from defects in materials and workmanship for a period of one (1) year from the date of purchase. MEDmetric® agrees within this warranty period to replace or repair, at MEDmetric’s® option and without cost to the purchaser, any part or parts, returned to MEDmetric®, with transportation charges prepaid, which upon MEDmetric’s® examination discloses to have been defective as a result of defects in materials and/or workmanship. In no event shall MEDmetric® be held liable for transportation expenses or any expense other than the cost of service, repair or replacement of defective  parts in accordance with provisions of this warranty. Warranty services may be arranged by contacting: 

MEDmetric® Corporation 

7542 Trade Street 

San Diego, California 92121-2412 

Telephone: (619) 536-9122 or (800) 995-6066; 

fax: (619) 536-9303 E-mail: info@medmetric.com

Misuse, abuse, neglect or unauthorized repairs or alterations shall void this warranty. 

No implied warranties of any nature whatsoever expressly including warranties of merchantability and fitness for a particular use, shall exceed the duration of this express warranty. It being recognized that the one (1) year period from the date of purchase is a reasonable warranty period for inspection of the products to allow determination as to whether there are defects in material and workmanship, and that the warranty period is a reasonable period for such an implied warranty to exist.  

MEDmetric® shall not be liable for incidental or consequential damages of any nature, and MEDmetric’s® sole liability shall be for the repair or replacement of defective parts pursuant to this warranty. 

MEDmetric® will not be liable for commitments or agreements made by any of its employees, agents or dealers not in compliance with this warranty. This warranty is expressly in lieu of any other warranty, express or implied. This warranty gives specific legal rights to the purchaser of MEDmetric’s® products covered hereby. Purchasers may also have legal rights which vary from state to state. 

Terms and Conditions of Sale

When MEDmetric® Corporation accepts an order, it shall become a binding contract subject to our terms and conditions of sale and will not be subject to countermand without the written consent of  MEDmetric® Corporation. 

Terms. All invoices are due and payable thirty (30) days from date of shipment unless otherwise indicated. 

Prices. Orders will be priced at the prices in effect at the date of the order. Prices are subject to change without notice. 

Deliveries. Promises of shipments are estimated as closely as possible. Our Traffic Department cooperates fully in obtaining the best service from the transportation companies, and we always use our best efforts to ship within the time promised. However, we cannot always guarantee delivery dates. Unless otherwise agreed, all merchandise is F.O.B. our factory. Special shipping requests are subject to additional charges. 

Returns. Merchandise may not be returned without our prior consent. 

Warranty. All products manufactured by MEDmetric® Corporation are warranted exclusively in accordance with our limited warranty. 

Product Changes. We reserve the right to redesign, alter or modify our products without incurring any liability whatsoever to any prior buyer of our products. 

Product Specifications. All products supplied by us are designed to perform in accordance with specifications that are available from us and should be consulted prior to use. If you are in doubt, please consult our Engineering Department. 

Attorney's Fees and Costs. If any acknowledgments at law or inequity is necessary to enforce or interpret the Terms of Purchase and Sale contract which becomes mutually binding upon our acceptance of your order, subject to the terms and conditions of sale herein, the prevailing party shall be entitled to reasonable attorneys fees, costs and necessary disbursements in addition to any other relieve to which he may be entitled. 

Appendix A

Side-to-Side Differences: ACL Deficit Patients Compared to Normal Subjects

	Anterior Displacement
	Normal Subjects (N=100)

Side to Side Difference (mm)
	ACL Deficit Patients (N=50)

Side to Side Difference (mm)

	
	Mean
	S.D.
	Range
	Mean
	S.D.
	Range

	30º/20 lb.
	0.5
	0.5
	0 to 2.0
	5.6
	2.2
	2 to 11

	30º/Max. Manual
	0.8
	0.8
	0 to 3.5
	6.3
	2.2
	3 to 11

	30º  Compliance 

Index
	0.1
	0.2
	0 to 1.0
	1.3
	0.9
	0 to 3.5

	30º Active Displacement
	0.9
	1.0
	0 to 4.0
	4.2
	2.4
	0 to 12


	

MILLIMETERS
Side-to-Side difference: 30º/20 lb. Anterior displacement
	

MILLIMETERS
Side-to-Side difference: manual maximum displacement



	

MILLIMETERS
Side-to-Side difference: compliance index
	

MILLIMETERS
Side-to-Side difference: knee extension active displacement


Test-Retest Variation

	

Total displacement variation versus side to side difference variations at 30º  of flexion and 20 lb.. of  force. Each subject was tested 5 times by a single examiner.
	

Test-retest variation distribution: side to side                                             difference in anterior displacement




Subject Number
Side -to- Side

Difference

Variation (mm)
	Test-retest variation of each subject's side to side difference (30º/20 lb. passive anterior displacement). Each subject was tested five times by a single examiner.




Measurement Accuracy (Cadaver Studies)
The design and configuration of the KT1000™ has been thoroughly tested in the research 



Histogram shows how the average absolute error was distributed over  a range of skeletal displacements from 1 to 10 mm. The average  absolute error for all 34 measurements was 0.39 mm (± 0.25 mm).  The magnitude of the absolute error was relatively independent of the  size of the 20° anterior drawer (Lachman test).

laboratory. The graph below shows that over a range of stability measurements from 1 mm to 10 mm, the average absolute error between the known, electronically monitored skeletal displacements and KT1000™ measurements was less than 0.5 mm.


KT1000™ readings against actual skeletal displacements as measured through Steinman pins placed in the femur and tibia. The correlation coefficient for the instrumented versus known skeletal displacement was very high (r=0.979). The small errors, due probably to soft tissue compliance, were randomly distributed about the true mean relationship                             between the instrumented and known tibial displacements.
Also tested, but not shown here, was the effect of an experimentally induced knee effusion on KT1000™ measurement accuracy. This was done in cadavers by injecting saline fluid under pressure during instrumented testing. The pressurized saline simulation of a tens hemarthrosis or effusion showed that the instrumented displacement measurement could be made to increase by 1 mm to 2 mm unless increased posterior pressure was maintained upon the patellar reference pad to keep the patella seated upon the femoral notch during the test.

Cadaver testing of the clinical knee ligament instrument for accuracy and repeatability was performed. The direct skeletal displacements in the cadaver leg were monitored through pins placed in the femur and the tibia. The statistical correlation coefficient (r=0.979) showed good agreement between instrument measurements and actual motions. The average arithmetic error was 0.16 mm (± 0.44 mm).

Appendix B



Electronic Specifications for the KT2000

Force: 1 pound of force = 10 mv

Displacement: 1 mm displacement = 25 mv

Interface Cord Color Code
1. Orange: Force

2. Black: Signal Common

3. Red: Displacement

4. Green: Pen down switch terminal for X-Y plotter

5. White: Pen down switch terminal for X-Y plotter

6. Silver: Shield



43

_1160568824.doc
�



�
















_1160568828.doc
�



�
















_1160568831

_1160568833

_1160568834.doc
�



�
















_1160568830.doc
�



�
















_1160568826.doc
�



�
















_1160568821

_1160568822.doc
�



�
















_1160568819.doc
�



�
















